| Evident FV40000 CLSM START-UP PROCEDURE |  |  |
| --- | --- | --- |
| 1) Start by switching on computer – it has to be on before the system starts. Login with the Windows account BICuser (pw=BICuser). |  |  |
| 2)If will be used, turn on climate control: CO2/O2 at the bottom, Temperature in the middle and Ventilator on the top. The switches are at the back. To start the heating and CO2, use arrows (the indicator should show ON). |  |  |
| 3) Turn on the green master switch on the top (3). Turn on the key to a ON position and pirullinos for the lasers that would be used (i.e. 405;488; 561;640 – all have separate controls) (4). |  |  |
| 5) Turn on the Evident pad (switch at the back) |  |  |
| 4) Open the internet browser, login into iLab and go to the Kiosk. Start your session. Wait until the status will become . Note that late arrival (+30min) is automatically considered as NO SHOW in iLab, so a new session needs to be created. |  |  |
| 5) Start the CellSense software. |  |  |
| 6) For the best imaging results, clean objectives you intend to use. |  |  |

| Evident FV40000 CLSM SHUTDOWN PROCEDURE |  |
| --- | --- |
| Always shut down the instrument at the end of your session unless the next user is standing behind you. |  |
| 1) If an oil or water objective is used, clean the objective with lens paper. Also clean the slide holder on the stage if this has been used. |  |
| 2) Switch to 2x objective. Lower the stage until the lower Z-limit reached. |  |
| 3) Close the CellSense software and finish your session in iLab Kiosk. |  |
| 4) Sign out of BICuser account. Shut down Windows (optional). |  |
| 5) Switch off the Evident pad (5) unit. |  |
| 6) Switch off the pirullinos for lasers followed by the Key on Stand ON position (4). Turn off the green master switch (3). |  |
| 7) If the climate control was used, turn all units off. |  |
| 8) Dispose of used lens paper and other items in the trash boxes situated at the entrance to the facility. |  |