| Zeiss 980-Airy START-UP PROCEDURE |  |  |
| --- | --- | --- |
| 1) Switch on the Power button (1) Turn the key for the lasers (2) |  |  |
| 2) Switch on the Components button (3) | 3 |  |
| 3) Switch on the software computer, and login with the Windows account BICuser (pw=BICuser). |  |  |
| 4) Open the internet browser, login into iLab and go to the Kiosk. Start your session. Wait until the status will become . Note that late arrival (+30min) is automatically considered as NO SHOW in iLab, so a new session needs to be created. |  |  |
| 5) Start the ZEN blue software . Say yes to stage calibration if asked. |  |  |
| 6) For the best imaging results, clean objectives you intend to use. |  |  |

| OBJECTIVE | NA | WD | Immersion | Coverglass thickness |
| --- | --- | --- | --- | --- |
| 5x | 0.16 | 18.5mm | dry | 0.17mm |
| 10x | 0.3 | 2.0mm | dry | 0.17mm |
| 20x | 0.8 | 0.55 mm | dry | 0.17mm |
| 40x W | 1.2 | 0.28mm | water | 0.14-0.19mm |
| 63x Oil | 1.4 | 0.19mm | oil | 0.17mm |

| Zeiss LSM980-Airy CLSM SHUTDOWN PROCEDURE |  |
| --- | --- |
| Always shut down the instrument at the end of your session unless the next user is standing behind you. |  |
| 1) If an oil or water objective is used, clean the objective with lens paper. Also clean the slide holder on the stage if this has been used. |  |
| 2) Switch to 5x objective. Lower the stage until the lower Z-limit reached. |  |
| 3) Close the ZEN software and finish your session in iLab Kiosk . |  |
| 4) Sign out of BICuser account. Shut down Windows. |  |
| 5) Switch off the Components (3). |  |
| 6) Turn the Laser Key (2) to off, followed by switching of the Main Power (1). |  |
| 7) Put on the dust protective cover. |  |
| 8) Dispose of used lens paper and other items in the trash boxes situated at the entrance to the facility. |  |