| NIKON CrEST SDM START-UP PROCEDURE |  |  |
| --- | --- | --- |
| 1) Optional for temperature + CO2 control, switch on the OKO lab temp control unit. It will take approx. 30min to warm up. |  |  |
| 2) Switch on the Kinetix camera, button (1) on left. Wait for the camera initialization to be finished when the LED (arrow) on the camera is switched off. |  |  |
| 3) Turn on the CrEST Spinning Disk Unit (2) by pressing the POWER button. | 2 |  |
| 4) Switch on the Celesta laser unit (3) (on the rack on the right) and turn on the laser key (4). | 4 3 |  |
| 5) Switch on the Nikon control unit (5) and the fluorescence P800 LED unit (6). | 5 6 |  |
| 6) Switch computer, and login with the Windows account BICuser (pw=BICuser). |  |  |
| 7) Open the internet browser, login into iLab and go to the Kiosk. Start your session. Wait until the status will become |  |  |
| 8) Start the NIS-elements software , login with your KI ID name (choose from drop-down menu) and password (same as username) |  |  |
| 9) For the best imaging results, clean objectives you intend to use. |  |  |

| OBJECTIVE | NA | WD (mm) | Emm. Iris D (%) |
| --- | --- | --- | --- |
| 4x | 0.2 | 20 | 100 |
| 10x | 0.45 | 4 | 90 |
| 20x | 0.8 | 0.8 | 80 |
| 20x LWD | 0.7 | 2.3 | 70 |
| 25x Sil | 1.05 | 0.55 | 84 |
| 60x Oil | 1.4 | 0.13 | 47 |
|  |  |  | D=2000*NA/mag |

Nikon CrEST SDM SHUTDOWN PROCEDURE

Always shut down the instrument at the end of your session unless the next user is standing behind you.

- If oil or silicon objective was used, clean the objective with lens paper!! Also clean the slide holder on the stage if this has been used.

- Take away the used stage insert and place them inside the incubator to the right side.

- Lower stage all the way down.

- Switch to 4x objective.

- Close the NIS-elements software and finish your session in iLab Kiosk.

- Sign out of BICuser account and shut down Windows / Computer.

- Switch off the Nikon control unit and the P800 LED unit.

- Turn the laser key to off, followed by switching of the Celesta laser unit

- Press the power button on the CrEST unit.

- Switch off the Kinetix camera

- Dispose of used lens paper and other items in the trash boxes situated in the wet lab at the entrance to the facility.